## Slide 1: Development of Thermal Energy Harvesting Power Supply with an Internal Startup Circuit for Pacemakers
Pre-Incubatee work report
Under the guidance of
Dr Kailash Chandra Ray
Professor-In-Charge(Incubation Centre, IIT Patna)
Submitted By : Yaman Parasher

## Slide 2: Working of the Proposed System
TEG(Vin)
`40-60mV
Ring Oscillator
Modified Charge Pump
Cpst(Capacitor)
Comparator 1
Vcmp1
Vth or Vref of comparator
is in between 1.5-2.4V
Steady State Boost Converter
Startup Boost Converter
Oscillator
MUX
From upper left to right followed by bottom
right to bottom left
Flow
Comparator 2
o/p Voltage
This modified CP helps to achieve 2.5V in 8s
Charge upto Vpst
Sets when Vpst reaches 2.5V
Vpst 2.5V for SUBC comparator
Provide clock signal to SUBC counter
Control circuit in it provide clock phases for both itself and SSBC
Cout
Capacitor
Cout rises Vout till it reaches Vth of comparator 2
Sets when Vout >Vth
o/p enable

## Slide 3: Systematic Design of Boost Converter associated with its Control Circuit

## Slide 4: Basic architecture of Boost Converter
In low power application, we generally avoid diode for relevant switching ,due it high forward voltage loss associated with it.
Therefore in our proposed system, we employ low power MOS of P & T type to cope up with the switching losses.
Working
At the vey beginning, the switch S1 is closed with S2 open.
Due to this ,the potential at the input end will develop or accumulate energy in the inductor due to flow of current .
In the next stage, when switch S1 is open and S2 is close, there comes a back emf which will result in greater accumulation of voltage in the circuit which is equal to v(back emf) and the one provided by the input end V(in).This whole thing will make the capacitor charge up to more voltage level at subsequent stages of clock cycle ultimately leading to the increment in the voltage at the output end.

## Slide 5: Boost Converter(CCT structure)
~40mV-60mV
Vout Expected to be 1-3V(regulated) that makes power of nearly 130µW with Vout=2.6V
1.5Ώ
400µH
400nF
Vref is taken to be somewhere between 2.5V
There has to be two level shifter, one from high to low(just after the comparator output) and the other one from low to high(just after the output of the decision circuit ,i.e. from the output of NOT and NAND gate giving switching power supply to both MOS above in the circuit.
Corresponding waveform for high and low power loads
A | B | AB’
0 | 0 | 1
0 | 1 | 1
1 | 0 | 1
1 | 1 | 0
NAND gate truth table(for reference)

## Slide 6: Case 1(When Vref > Vout)
1
A | B | AB’
0 | 0 | 1
0 | 1 | 1
1 | 0 | 1
1 | 1 | 0
pmos
nmos
=2.5V
1
1
0
0
~40mV-60mV
1.5Ώ
400µH
Here Vref is taken as 2.5V
400nF
Subcase 2 (When Counter gives logic‘0’)
0
In this case nmos act as open circuit ,while
pmos behave as a conducting wire
nmos behave as open circuit at logic 0 and conducting at logic 1
pmos behave as open circuit at logic 1 and conducting at logic 0

## Slide 7: Case 1(When Vref > Vout)
1
A | B | AB’
0 | 0 | 1
0 | 1 | 1
1 | 0 | 1
1 | 1 | 0
pmos
nmos
=2.5V
1
1
0
1
1
pmos behave as open circuit at logic 1 and conducting at logic 0
nmos behave as open circuit at logic 0 and conducting at logic 1
NAND gate truth table(for reference)
~40mV-60mV
1.5Ώ
400µH
Here Vref is taken as 2.5V
400nF
Subcase 1 (When Counter gives logic‘1’)
In this case nmos act as close circuit or conducting ,while
pmos behave as a open circuit.

## Slide 8: Analysis
ɸ1 | ɸ2 | Interpretation
1 | 1 | (M1 is closed, M2 is open) Accumulation of energy at inductor
0 | 0 | (M1 is open, M2 is closed) Transfer of energy from inductor to capacitor(charging of capacitor)
0 | 1 | (M1&M2 both remain open) Discharging of capacitor, building up potential at load

## Slide 9: Proposed TE energy harvesting architecture
~40mV-60mV
Preferably a ring oscillator due to low power consumption
MUX
Control circuit for both SUBC & SSBC
Helps in generating clock phases for SSBC and providing
Vout at the initial stage so that comparator 2 get set for function
Proposed ring oscillator
Proposed charge pump
Same Boost converter
Cpst(Capacitor)
SUBC
SSBC

## Slide 10: Ring Oscillator in the Proposed System
The output voltage of TEG is very low. To make the oscillator work with such a low voltage a ring oscillator is employed in the circuit .A general Ring
oscillator is made up of successive inverter based stages with its output buffer. A typical inverter is made by the combination of nmos and pmos. Thus in Order to make the oscillator work we need to reduce the MOS threshold voltage. The only way to do it is by Forward Body Biasing. In Forward Body biasing, the threshold voltage decreases by making the Vsb in the Nmos negative and positive in Pmos. It can be further explained by MOS threshold voltage equation.
Ring Oscillator(3-stage)
Output Buffer(chain of inverters)

## Slide 11: Charge Pump in the Proposed System
Modified Dickson Charge Pump
The output voltage of a conventional CP circuit after N stages is given by
Vout=Vin – Vth + N( Vph – Vth)
Where Vin is the input voltage, Vph is the clock amplitude and Vth is the threshold voltage of the MOS switches.
It could be clearly seen from the above equation that, in order to increase efficiency of Charge Pump, we need to minimize the threshold voltage.
It could be done by mainly by one process, which is the FBB(Forward Body Biasing) .
Conventional Charge Pump Circuit
In the modified version, body voltage of MOS switches of the odd(even) stages is being connected to the output voltage of the last even(odd) stage. This process increases Vsb which will ultimately result in the reduction of threshold voltage of MOS that is the Vth.

## Slide 12: Control Circuit of proposed TE harvesting architecture
MUX
Vcmp2
Vph1,SUBC
Vph2,SSBC
Vph2,SSBC
Vph2,SUBC
Towards
Analog circuit of Start Up Boost Converter
Note MUX is only used to provide clock phases at the initial stage to SSBC by SUBC, once SSBC gets the required Vout then it will generate phases for itself , thus eliminating need of SUBC
Towards
Analog circuit of SSBC
OE

## Slide 13: Schematic of MUX in Proposed System
Vph1,SSBC
Vph2,SSBC
Vph1,SUBC
Vph2,SUBC
VCMP2
OE
OE
OE
OE
Vph1,SSBC(OUT)
Vph2,SSBC(OUT)
Steady State Tri State Buffer
Steady State Tri State Buffer
Startup Tri State Buffer
Startup Tri State Buffer
A simple tristate is a simple buffer
,that has an input, output and a control (here OE).It provide output same as that of the input only when the control valve(OE) is set ‘HIGH’.

## Slide 14: Case1(When Vcmp2 is Set)
Vph1,SSBC
Vph2,SSBC
Vph1,SUBC
Vph2,SUBC
VCMP2
OE
OE
OE
OE
Vph1,SSBC(OUT)
Vph2,SSBC(OUT)
Steady State Tri State Buffer
Steady State Tri State Buffer
Startup Tri State Buffer
Startup Tri State Buffer
1
1
1
0
0
0
Vph1
Vph2
In this case,
SUBC Tristate buffer
won’t work
In this case, SSBC Tristate buffer will provide self generated clock phase to the output that will derive its circuit. Here in this case they wont need a SUBC to provide a clock phase

## Slide 15: Case1(When Vcmp2 is Reset)
Vph1,SSBC
Vph2,SSBC
Vph1,SUBC
Vph2,SUBC
VCMP2
OE
OE
OE
OE
Vph1,SSBC(OUT)
Vph2,SSBC(OUT)
Steady State Tri State Buffer
Steady State Tri State Buffer
Startup Tri State Buffer
Startup Tri State Buffer
0
0
0
1
1
1
Vph1,SUBC
Vph2,SUBC
In this case,
SUBC Tristate buffer
won’t work
In this case, SUBC Tristate buffer will provide clock phase to SSBC, that will drive its circuit.

## Slide 16: Potential Applications of our work

## Slide 17: THERMOELECTRIC GENERATOR FOR BIOMEDICAL IMPLANT
Convert thermal energy into electricity
Fig 2.1 Typical simplest thermoelectric generator
p-Type
n-Type
Heat Input
Heat rejected

## Slide 18: Thermo Electric Generator
Equivalent Circuit of TEG
Typical TEG

## Slide 19: Thermoelectric Module
Heat absorbed(cold side)
Electric Insulator(Ceramic)
Positive(+)
Negative(-)
Electrical Conductor(Copper)
P-Type
n-Type
Heat rejected(hot side)

## Slide 20: Points to Ponder in development of TEG
Typical power requirement for implantable medical devices range between 30 to 100µW.
Though device design and material of TEG can be optimized to increase device performance, the power output ultimately depends on the available temperature difference within the application.
This temperature differences are only readily available near the surface of the skin where the heat is emitted to the ambient environment.
The tissue near the skin surface can be modeled into three layers consisting of muscle, fat and epidermis(skin).
According to previous studies, temperature gradient of 1-5K is available in the fat layer, which may vary depending upon the body location, ambient environment and physical activities.
These conditions play a vital role in determining optimal thermoelectric device placement.
The temperature gradient is proportional to the fat thickness due to low thermal conductivity of fat.
That means –> An overweight person running without clothing will have the largest temperature gradient within the fat layer and as a result will provide the most power to a thermoelectric generator.

## Slide 21: Abstract
A complete thermal energy harvesting power supply for implantable pacemakers is presented in this presentation The designed power supply includes an internal startup and does not need any external reference voltage. The startup circuit includes a prestart up charge pump (CP) and a startup boost converter. The prestart up CP consists of an ultralow-voltage oscillator followed by a high-efficiency modified Dickson. Forward body biasing is used to effectively reduce the MOS threshold voltages as well as the supply voltage in oscillator and CP. The steady-state circuit includes a high-efficiency boost converter that utilizes a modified maximum power point tracking scheme. The system is designed so that no failure occurs under overload conditions.
Thus ,in the presentation a comprehensive analysis of the working of Boost converter circuit is done.